## Slide 1: Mass Spectrometry and Proteomics in Research
Kyle Swovick, PhD
Mass Spectrometry Resource Lab
URMC Center for Advanced Research Technologies
IND 532
3/21/2022

## Slide 2
URMC Mass Spectrometry Resource Laboratory, KMRB G-9828
(585) 275-0769
https://www.urmc.rochester.edu/research/mass-spectrometry.aspx
Sina Ghaemmaghami, PhD
Scientific Director,
Associate Professor (Biology)
(585) 275-4829
sghaemma@bio.rochester.edu
Kevin Welle
Technical Director
(585) 275-0769
Kevin_Welle@urmc.rochester.edu
Jennifer Hryhorenko
Technical Associate
(585) 275-0769
Jennifer_Hryhorenko@urmc.rochester.edu
Kyle Swovick, PhD
Director of Assay Development
(585) 275-0769
Kyle_Swovick@urmc.rochester.edu
Personnel
Notes: Our facility consists of four people, our Director, Sina Ghaemmaghami, who is also an associate professor for the Biology department. He is the master grant writer and experiment planning extraordinaire.
Kevin Welle is our technical director. He is the one that customers will actually interact the most with if they have questions regarding the technical side of the mass specs. He is also the main person in charge of performing small molecule experiments.
Jennifer Hryhorenko is our tech associate who does all of the sample prep for each proteomics experiment.
I am the newest member and I help Kevin with running and maintaining the instruments and help Jenny with sample prep in the event we are overwhelmed with samples. And then I am also in charge of public outreach and developing and implementing novel mass spec techniques.

## Slide 3
Proteomics and small molecule projects
Consultation
Sample preparation
Data analysis interpretation
Letters of support
Grant writing and publications
Services offered
Notes: And at the core, we offer a variety of services to fulfill pretty much any mass spec need you might have. We have the capabilities to perform both proteomic and small molecule projects. And we can help with a wide range of services; from helping you design experiments, preparing your samples, assisting with data analysis, writing letters of support, and then helping with both grant writing and publications.
But today, I’m just going to focus on how we use mass spectrometry for proteomics.

## Slide 4
LC-MS/MS for proteomics
“Proteomics represents the effort to establish the identities, quantities, structures, and biochemical and cellular functions of all proteins in an organism, organ, or organelle and how these properties vary in space, time, or physiological state”
National Research Council Steering Committee, Mol. Cell. Prot., 2002
“Proteins are not only the most common effectors of disease pathogenesis, but are also determinants of treatment response. Unfortunately, protein analysis is technically challenging, inefficient, and represents a major bottleneck in efforts to advance our understanding of human health and to develop effective therapeutics”
Cayer D et al., Human Molecular Genetics, 2016
Notes: So what exactly is LC-MS/MS based proteomics?
Shortly after the National Academy of Sciences said proteomics was the technology to develop in the post-genomics world in 2002, a group of pioneers in the proteomics world said *1st quote here*.
And a more recent perspective states *2nd quote here*
So e can see that there is an increasing interest in characterizing and quantifying proteins.

## Slide 5
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
So why measure at the protein pevel?
Notes: But why is it so important to measure at the protein level?
Does anyone have any thoughts?

## Slide 6
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Protein turnover
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
So why measure at the protein pevel?
Notes: Proteins are the what fulfill actual biological function

## Slide 7
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Protein turnover
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
So why measure at the protein pevel?
Notes: And just using genomics or transcriptomics isn’t sufficient because there are many differences between DNA, RNA, and protein

## Slide 8
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Protein synthesis and degradation
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
So why measure at the protein pevel?
Notes: I’ve only included a few examples here, but of course there are many others

## Slide 9
So why measure at the protein pevel?
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Protein synthesis and degradation
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
Notes: So measuring just measuring RNA or DNA or both doesn’t gives us information on a variety of things.
Anyone have any thoughts on what these can be?

## Slide 10
So why measure at the protein pevel?
Proteins are what do the real work
There are many disconnects between DNA, RNA, and protein
        Alternative splicing
        Nonsense mediated decay
        Post-translational modifications
        Protein synthesis and degradation
        Measuring just nucleic acids does not inform us on:
          Where the protein is
          How much of the protein there is
          How rapidly that protein is turned over (my specialty)
          Presence and location of post-translational modifications
Notes: The localization of a protein, how much actual protein there is, how rapidly that protein is turned over, or is a protein is modified and where on the protein it’s modified
Despite all of these benefits to measuring at the protein level, there are some serious difficulties to actually measuring protein levels, even with mass spec. And this is for a myriad of reasons

## Slide 11
Why is measuring protein levels so difficult?
Proteins have a wide range of biophysical properties
Puthenveeti and Vinogradova, JBC Reviews, 2019
Tang et al., J. Voice, 2017
William Eaton, Journal of Biological Physics, 2021
Membrane proteins
Pro-collagen
Hemoglobin
Notes: One of which is that proteins have a huge range of biophysical properties. You can have a whole bunch of membrane proteins that all interact with membranes in various ways, you can have proteins that form quaternary structures with other protein subunits like collagen and hemoglobin.

## Slide 12
Why is measuring protein levels so difficult?
Protein expression levels span a huge dynamic range
protein rank
[protein]
40 g/L
400 pg/L
https://www.proteinatlas.org/humanproteome/blood+protein/proteins+detected+by+immunoassay
Notes: Additionally, protein expression levels span a huge dynamic range. Here is a chart showing the levels of 420 proteins in human blood. The most abundant protein, albumin, is at 40g per liter while interferon gamma is only present at 400 picograms per liter.

## Slide 13
Why is measuring protein levels so difficult?
Protein expression levels span a huge dynamic range
protein rank
[protein]
40 g/L
400 pg/L
https://www.proteinatlas.org/humanproteome/blood+protein/proteins+detected+by+immunoassay
12 orders of magnitude!
Notes: That’s a difference in 12 orders of magnitude!

## Slide 14
Why is measuring protein levels so difficult?
Protein expression levels span a huge dynamic range
12 orders of magnitude!
https://www.physics.upenn.edu/~shethrk/courses/ast06/lectures/lecture2/s1.8.html
Notes: And just to put that in perspective, the difference in size between the average human height and the size of an atom is 11 orders of magnitude. So whatever method you use ideally can accurately measure proteins levels across that large of a span.

## Slide 15
Why is measuring protein levels so difficult?
There is no amplification (e.g. PCR)
Proteins are sticky
        Proteins are very common laboratory contaminants (keratin)
        Every protein is different
          Digesting proteins into peptides helps a lot
Notes: Additionally, there are some other factors.
For one, like how RNA-seq or genomics use PCR, we cannot amplify proteins so we are stuck with the amount we can extract.
Proteins are incredibly sticky. So they will pretty much stick to any surface they are in (tubes, pipettes, etc).
Also proteins are an incredibly common contaminant. No matter what, we will always see keratin in our mass spec data.
And then, similar to what I said earlier, every protein is different so no single preparation method suits all proteins. But digesting proteins into peptides (I’ll be discussing this later) helps a lot with that

## Slide 16
Why use LC-MS/MS to measure protein levels
Methods commonly used to measure protein expression levels:
Notes: But why use mass spectrometry to measure protein levels when we have many incredibly well established methods to do that?
I came up with a couple common techniques. Can any of you think of some common techniques?

## Slide 17
Why use LC-MS/MS to measure protein levels
Methods commonly used to measure protein expression levels:
Western blot
Notes: Yeah Western blot is probably the most famous

## Slide 18
Methods commonly used to measure protein expression levels:
Western blot
                Very established and well accepted
                Low-throughput
RNA-seq
Why use LC-MS/MS to measure protein levels
Notes: I debated on including this one, but a lot of people use RNA-seq data as a proxy for protein expression levels.

## Slide 19
Methods commonly used to measure protein expression levels:
Western blot
                Very established and well accepted
                Low-throughput
RNA-seq
                High-throughput
Why use LC-MS/MS to measure protein levels
Notes: Western blots are great because they are very well established, relatively cheap to perform, and are very well accepted.
However, when you start to have to analyze a bunch of samples, it can start becoming very expensive and time consuming. Plus you have to worry about non-specific binding, especially if you are studying post translational modifications.

## Slide 20
Methods commonly used to measure protein expression levels:
Western blot
                Very established and well accepted
                Low-throughput
RNA-seq
                High-throughput
                Very poor at predicting protein
                levels
Why use LC-MS/MS to measure protein levels
Notes: RNA-seq on the other hand is very high-throughput. But it can be expensive. And probably the biggest problem is that it is very poor at actually predicting protein levels.

## Slide 21
Methods commonly used to measure protein expression levels:
Western blot
                Very established and well accepted
                Low-throughput
RNA-seq
                High-throughput
                Very poor at predicting protein
                levels
Schwanhausser et al., Nature, 2011
Why use LC-MS/MS to measure protein levels
Notes: As shown here when mRNA copy number and protein copy number per cell are compared, there is a very very weak correlation. So if you want to see if a gene is unregulated, RNA-seq may not be the best method.

## Slide 22
Very sensitive (fmol)
Very precise (<1 dalton)
Can analyze complex mixtures
Sample is used up
Information poor (molecular weight only)
The power of mass spectrometry….
….and it’s shortcomings
Notes: So mass spec has some great advantages. One, it is very sensitive and can measure down to the femptomolar range. And is very precise (it can differentiate differences in mass that are less than 1 dalton).
Additionally it is able to measure thousands of proteins from complex mixtures like tissue.
But some shortcomings include that is uses up the sample so you can’t perform subsequent experiments on that sample.
Also, the data is relatively information poor (all it gives you is molecular weight)

## Slide 23
Determine the mass of analytes
Quantify the amount of a small molecule (i.e. drug or metabolite) in a biological matrix
Identify a protein(s) from a gel, blood, lysate, etc.
To quantify protein expression level changes between biological conditions
What can mass spectrometry be used for?
Notes: But what can mass spectrometry be used for? A lot really.
It can determine of course the mass of analytes.
But it is also able to quantify the amount of a small molecule (like a drug or metabolite) in a biological matrix.
Identify proteins from a gel, blood, lysate, and many other things.
And we can use it to quantify differences in protein expression levels in biological conditions.

## Slide 24
Chart: Year
| 2000 | 2001 | 2002 | 2005 | 2008 | 2011 | 2014 | 2017 | 2020 | 2021
Region 1 | 463 | 1104 | 1743 | 4554 | 6577 | 8404 | 10196 | 11161 | 14489 | 16121
Proteomics- a growing field
Notes: And it seems people are quickly starting to realize just how useful mass spec-based proteomics can be. Here I’m just plotting the number of articles in pubmed with “Proteomics” in the name since 2000. And it’s really rapidly growing.

## Slide 25: Increase in proteome coverage over the years
Chart: Publication Date
| 2006 | 2008 | 2012 | 2015 | 2018
Region 1 | 1000 | 1900 | 4500 | 7200 | 10217
Notes: And over the past few years, the technology in our field has only gotten better. In fact, in 2018, a paper was published that was able to identify nearly 11,000 proteins in one HeLa digest.

## Slide 26
Q-TOF
Triple
Quadrupole
MALDI-TOF
Technique to measure the mass to charge ratio of ions
What is mass spectrometry?
Q-TOF
Triple Quadrapole
MALDI-TOF
Orbitrap Tribrids and Hybrids
Notes: So I’ve been talking all about mass spectrometry, but what is it?
It’s really just a technique to measure the mass to charge ratio of ions. And there are a ton of different instruments that can do this, all working slightly differently.
The most common types of orbitrap tribrids and hybrids, Q-TOF (stands for quadrapole time of flight), a triple quadrapole (you’re frequently hear people refer to these as triple quads), and a MALDI-TOF(stands for matrix assisted laser desorption/ionization time of flight).

## Slide 27
What is mass spectrometry?
Q-TOF
Triple
Quadrupole
MALDI-TOF
Orbitrap Tribrids and Hybrids
Q-TOF
Triple Quadrapole
MALDI-TOF
Technique to measure the mass to charge ratio of ions
Notes: And in our facility, we use oribtrap instruments.

## Slide 28
Mass spectrometer design
Notes: While they may differ slightly in how they work, they all share the same layout.
First you inject a sample through an inlet.
Then it becomes ionized by an ion source.
Then the ion proceeds to a mass filter, which removes any undesired ions.
The ion will then go to the mass analyzer, which actually measures the mass.
And all of this needs to be performed under vacuum and is controlled through computer software.
Next I’m going to break each of these sections down in more detail.

## Slide 29
Inlet Source- HPLC
Separates molecules based off biophysical attribute (commonly hydrophobicity)
          Hydrophobicity of solvent increases as a function of time -> more hydrophobic molecules will elute later
Notes: So the inlet source is what feeds your sample into the mass spec. For most proteomics experiments, this is through an HPLC.
In proteomics experiments, HPLCs are used to separate peptides based off some biophysical property (most commonly hydrophobicity).
So over the course of a mass spec experiment, the hydrophobicity of the solvent running through the HPLC column increases so more hydrophobic peptides elute later in the experiment.

## Slide 30
Mass spectrometer design
Notes: So the peptides have left the HPLC now they go to the ion source.

## Slide 31
+ +
+
++ ++
+ +
+
++ ++
++
+
+
++
+
+
++
++ ++
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+ + +
+
+
+
+
+
Charged droplets
Capillary column
Mass Spectrometer
+2 kV to +5 kV
Solvated sample (from HPLC)
Desolvated ions
Heated desolation region
Ion Source- Electrospray Ionization (ESI)
Notes: Ionization is one of the most critical steps in protein based mass spectrometry. And actually, before electrospray ionization was invented in 1995, ionization methods were too harsh for peptides so protein-based mass spectrometry was nearly impossible. (The inventor of ESI was awarded the Nobel prize for it in 2002)
So as you can probably now guess, we ionize samples through electro spray ionization. In this method, the sample in liquid (from the HPLC) is injected into a column that has a voltage applied to it. At the end of this column, the liquid is sprayed into the mass spec. At this point, you have small droplets that contain the peptides that have now have a charge to them. They are then heated to 255C which will evaporate all of the moisture in the droplets leaving you with only dried ions at the end. At which point, they are focused into the mass spec

## Slide 32
Mass spectrometer design
Notes: So now our peptides have been ionized and from this point on, they are referred to as ions. The ions then go through a mass filter.

## Slide 33
Mass filter- quadrupole (Q)
Notes: One of the most commonly employed mass filter is a quadrapole.

## Slide 34
Mass filter- quadrupole (Q)
Mass Analyzer
Notes: Differing currents are run through each of the poles and like I said earlier, the ions come from the ESI through the quadruple, and then exit into the mass analyzer.

## Slide 35
Mass filter- quadrupole (Q)
Mass Analyzer
Notes: But let’s see what’s happening inside the quadrapole. With the different currents being run through the quadrapole, each ion oscillates through the poles and it eventually exits into the detector.

## Slide 36
stable path
unstable path
Mass filter- quadrupole (Q)
Mass Analyzer
Notes: However, not every ion is going to make it through. Some ions will not have a stable path through the poles. When this happens, the ion will hit one of the four poles and will not travel any further. But if it does have a stable path, the ions can indeed travel through.

## Slide 37
ΔCurrent = x amp
Scanning with quadrupole mass filters
Notes: The nice thing with these quadruples is that you can change the current being applied through them.
So let’s imagine we have three separate ions shown here by three differently colored amplitudes.
For this particular current, only the ion with blue amplitude is able to travel through the quadrapole.

## Slide 38
ΔCurrent = x amp
ΔCurrent = y amp
ΔCurrent = z amp
Scanning with quadrupole mass filters
Notes: But if we apply different currents, we can get both the orange and purple ions to travel through.

## Slide 39
m/z
time (∆ current)
ΔCurrent = x amp
ΔCurrent = y amp
ΔCurrent = z amp
Scanning with quadrupole mass filters
Notes: So we can change the current over time so that only particular ions enter the mass analyzer analyzer.

## Slide 40
Mass spectrometer design
Notes: So we have filtered out an unwanted ions, and now they will go into the mass analyzer

## Slide 41
Mass analyzer- Orbitrap
Zubarev and Makarov, Anal. Chem., 2011
Notes: Today I’ll just describe the mass analyzer that both our instruments use, the orbitrap mass analyzer. Although it is incredibly small (As shown here, the newest version is slightly larger than a nickel), it is incredibly powerful and has ushered in another boom in proteomic research

## Slide 42
Mass analyzer- Orbitrap
Savaryn et al., Proteomics, 2016
Notes: The basic principle behind an orbritrap is you have a magnetic rod that emits an electromagnetic field. Ions are fed into this field, where they spin around the rod.

## Slide 43
Mass analyzer- Orbitrap
Savaryn et al., Proteomics, 2016
Notes: The electrodes then read the oscillatory motion of the ions, creating a transient. And the transient frequency is unique for each m/z

## Slide 44
Mass analyzer- Orbitrap
Savaryn et al., Proteomics, 2016
Notes: And then a fourier transformation is performed

## Slide 45
Mass analyzer- Orbitrap
Savaryn et al., Proteomics, 2016
Notes: To give us a mass spectrum.
There are several major advantages to orbitrap mass analyzers, including:
They have ridiculous resolution (meaning they are better able to differentiate between isotopes)
Very high accuracy
Very high sensitivity

## Slide 46
Mass spectrometer design
Notes: And like I said earlier, all of these processes need to be performed under vacuum

## Slide 47
Mass spectrometer design
Notes: And all of these processes all controlled through some kind of software and data stored within a computer

## Slide 48: Schematic of a QExactive hybrid mass spectrometer
Notes: So with the basic layout, I’ll show you the layout of one of our mass spec, a QExactive hybrid. So peptides elute off the column and then ionized by this spray source. Then the ions are focused through all these lenses. Ions are then shot into a quadrapole mass filter and then into what’s called a C-trap. All this does is hold ions in pace until we have enough to inject them into the mass analyzer. Once there are enough ions, they go into orbitrap.
And then is we want to get the sequence of the ion (I’ll describe this more in depth later), a new set of ions are put into the c trap. This time instead of going into the orbitrap, they go into a collision cell where the ions are busted apart. Then they go back into the C trap to then into the orbitrap to measure the mass of the new ions. And these is repeated every 2.5 seconds

## Slide 49
Bottom-up proteomics (most common and what we do)
        Digest complex protein mixtures into peptides
        Obtain the fragmentation data of the peptides
        Match the peptide fragment spectra to a predicted spectra from a peptide amino acid sequence
        Infer the protein from the peptide sequence
        Top-down proteomics (very difficult)
          Measure the intact mass of a protein
          Fragment the protein
          Identify the protein from the fragmentation pattern
Top-down vs. Bottom-up proteomics
Notes: So now that we have a general idea how a mass spec works, let’s dig a little deeper into proteomics.
Within proteomics, there are two general fields; bottom-up and top-down proteomics

## Slide 50
Bottom-up proteomics (most common and what we do)
        Digest complex protein mixtures into peptides
        Obtain the fragmentation data of the peptides
        Match the peptide fragment spectra to a predicted spectra from a peptide amino acid sequence
        Infer the protein from the peptide sequence
        Top-down proteomics (very difficult)
          Measure the intact mass of a protein
          Fragment the protein
          Identify the protein from the fragmentation pattern
Top-down vs. Bottom-up proteomics
Notes: Bottom up proteomics is by far the most common and is what we do, while top down is very challenging

## Slide 51
Bottom-up proteomics (most common and what we do)
        Digest complex protein mixtures into peptides
        Obtain the fragmentation data of the peptides
        Match the peptide fragment spectra to a predicted spectra from a an amino acid sequence
        Infer the protein from the peptide sequence
        Top-down proteomics (very difficult)
          Measure the intact mass of a protein
          Fragment the protein
          Identify the protein from the fragmentation pattern
Top-down vs. Bottom-up proteomics
Notes: Just to briefly describe them though, bottom-up proteomics is where we digest a complex mixture into peptides, obtain the fragmentation data of the peptides, match the peptide fragment spectra to a predicted one and then infer the presence of protein from the peptide sequence

## Slide 52
Bottom-up proteomics (most common and what we do)
        Digest complex protein mixtures into peptides
        Obtain the fragmentation data of the peptides
        Match the peptide fragment spectra to a predicted spectra from a peptide amino acid sequence
        Infer the protein from the peptide sequence
        Top-down proteomics (very difficult)
          Measure the intact mass of a protein
          Fragment the protein
          Identify the protein from the fragmentation pattern
Top-down vs. Bottom-up proteomics
Notes: While top-down proteomics measures the mass of an intact protein. And then you fragment that protein and verify the identity of the protein from the fragmentation pattern.

## Slide 53
Routine bottom-up proteomics workflow
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Notes: Since we only do bottom-up proteomics, I’ll describe a typical workflow for an average experiment.
First, we extract proteins from a biological matrix

## Slide 54
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: Then we digest those proteins into peptides

## Slide 55
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: Then optionally, we can divide those peptides into numerous fractions based on some biophysical property

## Slide 56
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: Now we load the peptides onto the mass spec (well HPLC really) and they are separated using reverse phase chromatography

## Slide 57
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: The peptides are ionized, making them go from the liquid to the gas phase

## Slide 58
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: Then we measure the mass of the peptides with a mass spectrometer

## Slide 59
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: Then isolate and fragment the peptides in the mass spectrometer

## Slide 60
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: We then measure the mass of those fragments using the mass spectrometer

## Slide 61
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: We then use this data to identify the peptide using a computer software

## Slide 62
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: We can then quantify the peptides using intensity data.

## Slide 63
Extract proteins from biological matrix
Digest proteins into peptides
(Optional) Fractionate peptides based on some biophysical property
Separate peptides using reverse phase chromatography
Ionize peptides (liquid to gas phase)
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Use this data to identify the peptide using computer software
Quantify peptides using intensity data
Statistics and differential expression
Routine bottom-up proteomics workflow
Notes: And then perform statistics and measure differential expression if that’s what we’re interested in.
That was a whirlwind of steps, so now I’m going to break one down and describe each one in more detail.

## Slide 64: 1) Extract protein from a biological matrix
Every sample and tissue is different
General method:
        Lyse cells using a detergent (SDS, RIPA) or urea followed by mechanical disruption (sonication, bead beating, liquid nitrogen grinding)
Notes: So the first step is to extract protein from some type of matrix. And of course, I made that sound incredibly simple, but the method for doing this can vary dramatically for every type of sample and tissue and what the goal of the experiment is.
But the general method is cells are lysed using a detergent or are based buffer followed by some kind of mechanical disruption which can include sonication, bed beating, and liquid nitrogen grinding.

## Slide 65: 2) Digest proteins into peptides
Almost always use enzymes (Trypsin by far the most common followed by Lys-C)
Notes: So now that we have extracted our proteins, we want to break them down into peptides, where now each molecule will behave more similarly to the next.
To do this, we almost always use enzymes (we use trypsin about 99% of the time).

## Slide 66: 2) Digest proteins into peptides
Almost always use enzymes (Trypsin by far the most common followed by Lys-C)
Notes: The general way this works is you reduce the disulfide bridges so now your proteins are unfolded

## Slide 67: 2) Digest proteins into peptides
Almost always use enzymes (Trypsin by far the most common followed by Lys-C)
Notes: And then the enzymes will cut at very specific locations (in the case of trypsin, the carboxyl side of lysine and arginine)

## Slide 68: 2) Digest proteins into peptides
Almost always use enzymes (Trypsin by far the most common followed by Lys-C)
Notes: So now we have a bunch of smaller peptides.

## Slide 69: 4) Separate peptides using reverse phase chromatography
High Performance Liquid Chromatography (HPLC)
(The Robin To Mass Spec’s Batman!)
By connecting a C18 column to the HPLC, we are able to separate peptides and significantly reduce the complexity of the sample entering the mass spec at any given time
Notes: We then load our peptides onto the HPLC where they can be separated using reverse phase chromatography. Typically C18 columns are connected to the HPLC so we are able to separate peptides and significantly reduce the complexity of the sample entering the mass spec at any given time.

## Slide 70
Bead
C18 chains
C18 resin
Notes: To just visualize what the C18 and HPLC are doing, we’ll look at it a little closer. So the C18 resin is just a bunch of beads with C18 chains on them.

## Slide 71
Bead
C18 chains
C18 resin
Notes: And when we load peptides onto them, they will become stuck to the C18

## Slide 72
Column packed with C18 beads
To Mass Spec
From HPLC
Peptide separation using C18 columns
Notes: So let’s look at the interplay between HPLC and mass spec. Here we have peptides coming from the HPLC into a column that is packed with that C18 resin, which is connected to the mass spec.

## Slide 73
Peptide separation using C18 columns
Peptides bind due to the interactions between the C18 and amino acid side chains
To Mass Spec
From HPLC
Notes: When we flow peptides through the HPLC, the amino acid side chain will bind to the C18

## Slide 74
To Mass Spec
From HPLC
The HPLC pumps an increasing amount of organic solvent through the column
Hydrophilic peptides move through the column faster than hydrophobic ones
Peptide separation using C18 columns
Notes: And then the HPLC pumps an increasing amount of organic solvent through the column.
Because of this, hydrophilic peptides will move through the column faster than hydrophobic peptides.

## Slide 75
The HPLC pumps an increasing amount of organic solvent through the column
Hydrophilic peptides move through the column faster than hydrophobic ones
Intensity
Retention Time
To Mass Spec
From HPLC
Peptide separation using C18 columns
Notes: So each peptide has a unique time that it is retained on the column before it goes into the mass spec. And we can plot that as the retention time.

## Slide 76: Mass spectrometer data acquisition
Measure the mass of peptides using a mass spectrometer
Isolate and fragment the peptides in the mass spectrometer
Measure the mass of the fragments
Notes: Then the peptides are ionized and go into the mass spec where their mass is measured, fragmented, and the mass of those fragments in measured.
Let’s look at these steps a little closer.

## Slide 77: Mass spectrometer data acquisition
Data-Dependent Acquisition (DDA)
Notes: So I will describe how data is acquired through a method called data-dependent acquisition.
So for any given retention time (remember the time a set of peptides elute off the C18, the mass spec will measure the mass of all the peptides with a certain mass range. Here I’m showing 400-1000 m/z.

## Slide 78: Mass spectrometer data acquisition (DDA)
“MS1” Scan
“Precursor” Scan
“Full” Scan
Notes: And to see a simplified version of the spectra, here is what we typically see. A bunch of lines at different m/z and each with different height (called intensity), roughly corresponding to their abundance.
These scans are commonly called MS1, precursor, or full scans.

## Slide 79: Mass spectrometer data acquisition (DDA)
Notes: So after an MS1 scan is done, the mass spec will pick one of the peptides from the MS1 scan and perform an MS2 scan on it

## Slide 80: Mass spectrometer data acquisition (DDA)
Notes: And it will do a whole bunch of different MS2 scans throughout the MS1 mass range.

## Slide 81: Mass spectrometer data acquisition (DDA)
“MS2” Scan
“Fragment” Scan
Notes: Let’s look at little closer at these MS2, or fragment scans. They look pretty similar to the MS1 scans

## Slide 82: Mass spectrometer data acquisition (DDA)
Notes: But let’s also just look at it from the spectra level. The instrument did an MS1 scan and found this peptide at 850 m/z and wanted to sequence that. So it fragmented just that one peptide into smaller bits and performed another scan (an MS2 scan) on all those tiny bits measure their mass. But what are all these b and y ions I have listed you ask and what exactly is happening here you ask? Let me explain

## Slide 83
Intensity
m/z
100
200
300
400
K+1
88.01
KY+1
201.12
E+1
90.06
LE+1
204.10
YLE+1
317.18
Intensity
m/z
404
405
KYLE+1
404.21
collision induced dissociation (CID)
KYL+1
315.16
Tandem MS (MS/MS)
Notes: So let’s look at an example case of a fake peptide called KYLE. In the MS1 scan, it has an m/z of 404.21 and is in the +1 charge state. Honestly, there are a ton of other peptides that can have this same m/z so it could be KYLE or it could be something else. The mass spec has decided it is interested in knowing if this peptide is actually KYLE.

## Slide 84
Intensity
m/z
100
200
300
400
K+1
88.01
KY+1
201.12
E+1
90.06
LE+1
204.10
YLE+1
317.18
Intensity
m/z
404
405
KYLE+1
404.21
collision induced dissociation (CID)
KYL+1
315.16
Tandem MS (MS/MS)
Notes: What the mass spec does is it selects just the 404.21 peptide and slams it really fast into an inert gas in a process called collision induced dissociation.

## Slide 85
Intensity
m/z
100
200
300
400
K+1
88.01
KY+1
201.12
E+1
90.06
LE+1
204.10
YLE+1
317.18
Intensity
m/z
404
405
KYLE+1
404.21
collision induced dissociation (CID)
KYL+1
315.16
Tandem MS (MS/MS)
b1
b2
b3
y3
y2
y1
Notes: What this does is it breaks the peptide at every possible peptide bond so we now a ton of different fragments (commonly referred to as fragment ions) called b and y ions. And that just corresponds to the location of the fragment ion with respect to the precursor ion

## Slide 86
Intensity
m/z
100
200
300
400
K+1
88.01
KY+1
201.12
E+1
90.06
LE+1
204.10
YLE+1
317.18
Intensity
m/z
404
405
KYLE+1
404.21
collision induced dissociation (CID)
KYL+1
315.16
Tandem MS (MS/MS)
b1
b2
b3
y3
y2
y1
Notes: So the first amino acid on the n terminus is b1 and then as you continue to the c terminus, you increase b by 1. And then from the c terminus, you start at y 1 and then increase y by one for every amino acid along to the n terminus

## Slide 87
Intensity
m/z
100
200
300
400
K+1
88.01
KY+1
201.12
E+1
90.06
LE+1
204.10
YLE+1
317.18
Intensity
m/z
404
405
KYLE+1
404.21
collision induced dissociation (CID)
KYL+1
315.16
Tandem MS (MS/MS)
Notes: But we can now figure out the amino acid composition of each fragment ions. And the idea is very similar to the idea of DNA shotgun sequencing.

## Slide 88
Intensity
m/z
100
200
300
400
1
S+1
88.0
S
2
I+1
01.12
SIN+1
315.16
404.21
E+1
90.06
LE+1
204.10
YLE+1
317.18
114.04 = L
113.08 = Y
87.03 = K
KYLE+1
404.21
90.06 = E
Determination of a peptide sequence
from an MS/MS Spectra
Notes: We now have essentially a ladder of increasing molecular weight ions. And we know the mass of each amino acid. So we can figure out the difference in mass between two fragments and deduce what the amino acid is.
And this is commonly referred to as MS/MS (because we do two MS scans) or tandem mass spec

## Slide 89: Mass spectrometer data acquisition (DDA)
Notes: So now we know the general way that this happens.

## Slide 90: Mass spectrometer data acquisition (DDA)
Notes: But within one MS1 scan, the instrument will select several peptides to fragment for MS2 scans. As shown here, each peptide is highlighted in a different color, and then we receive different colored MS2 scans that correspond to each specific peptide from the MS1.

## Slide 91: 7) Isolate and Fragment the Peptides in the Mass Spectrometer
m/z
ΔCurrent = x amp
ΔCurrent = y amp
ΔCurrent = z amp
Notes: But how is the mass spec able to pick specific peptides to fragment? Remember how I said the quadrapole mass filter can remove any ions you don’t want? Well that feature is incredibly useful now. In this example, the mass spec found three peptides in the MS1 scan, the blue, orange, and purple peptides. We can now apply very specific current to the quadrapole so that we can get only the blue, orange, or purple peptides and then we can fragment JUST that one peptide.

## Slide 92: Mass spectrometer data acquisition (DDA)
Notes: And then the mass spec just does this over and over and over again every 2.5 to 3 seconds.

## Slide 93: Mass Spectrometer Data Acquisition- Real Results
MS1 Scan
Notes: Now that we understand how data is collected, let’s look at an actual spectra. So this is an MS1 scan from one particular retention time (37.83 minutes). You can see that there are a ton of peptides here that could be interesting to sequence.

## Slide 94: Mass Spectrometer Data Acquisition- Real Results
“MS2” Scans
635.3434
Notes: So after this, the mass spec selects specific precursors. In this case, it picked the 635.3434 m/z precursor and fragmented it, giving us this MS2 scan.

## Slide 95: Mass Spectrometer Data Acquisition- Real Results
“MS2” Scans
473.8944
473.8944
Notes: And then after that it fragmented 473.8944.

## Slide 96: Mass Spectrometer Data Acquisition- Real Results
Notes: And then did 18 more MS2 scans before doing another MS1 scan.

## Slide 97: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: Okay now we have MS2 spectra but now we have to search this data. First, you take a known database, or FASTA file, of protein sequences corresponding to whatever organism you are searching.

## Slide 98: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: And then in silico digest those proteins into peptides

## Slide 99: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: And because these peptides behave fairly regularly, we can create theoretical spectra for each peptide

## Slide 100: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: We take our experimentally generated MS2 spectra

## Slide 101: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: And compare it to the in silicon generated spectra and score that relationship

## Slide 102: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: If that score is low, the search algorithm knows that this particular MS2 spectrum doesn’t correspond to that specific in silico generated peptide.

## Slide 103: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: So it will then go to the next spectra.

## Slide 104: 9) Identify the Peptide Using Computer Software
Modified from Kong et al., Nat. Methods, 2017
Notes: And for this pair, the spectra have a high level of similarity, so they received a high score, indicating that this MS2 spectra came from the fragment ion AFLIEEQK. And this process is repeated over and over and over for every MS2 scan.
And then we infer the protein presence from the peptide presence and perform various statical analyses on the proteins.

## Slide 105
Real-world applications of LC-MS/MS
Notes: So that’s the very general workflow for how we do a proteomics experiment. Now I want to show you all some of the things that you can do with proteomics.

## Slide 106: Example Application of LC-MS/MS: Identification of Interacting Proteins (IP-MS)
Notes: One incredibly useful scenario for this type of identification is through identifying proteins that are interacting with a protein of interest through immunoprecipitation mass spectrometry or IP-MS.
Let’s say your protein of interest is the red dot and you what to know what these colored proteins that interact with it are. You can purify your proteins through a Co-IP and then we run the IP on the mass spec and we can identify whatever proteins Co-IP with your protein.

## Slide 107: Example Application of LC-MS/MS: Identification of Interacting Proteins (BioID)
Sears et al., Methods Mol. Bio., 2020
Notes: In a similar fashion, a more recent technology that’s becoming more popular is BioID. In this method, you create a fusion protein of your favorite protein. This fusion protein can then biotinylate any other protein that interacts with your protein in the cell. You can then isolate any biotinylated proteins with strepavidin and then we can identify whatever proteins were biotinylated.

## Slide 108: Example Application of LC-MS/MS: Identification of Interacting Proteins (IP-MS)
Lentini et al., Nat. Comm., 2020
Notes: As an example of customer data that became a publication, the Fu lab in the Biology department on river campus, wanted to know what proteins were interacting with METTL2A, METTL2B.
They performed an IP using wither METTL2A or METTL2B as bait, ran that on a gel, and when they silver stained it, they saw unique proteins at the black arrows. So they cut those bands out and gave them to us to perform mass spec on and found that they not only interact with each other but a novel interaction partner in DALDR3 and where able to perform various biochemical assays characterizing the nature of this interaction.

## Slide 109: Example Application of LC-MS/MS: Measuring protein degradation kinetics
Swovick et al., Mol. Cell. Prot., 2021
Notes: I might be a little biased, but another really cool use of proteomics is measuring the dynamics of protein degradation. By using isotopically labeled lysine and arginine, we are able to calculate the rate at which every protein in the protein in the proteome is being degraded

## Slide 110: Example Application of LC-MS/MS: Measuring protein degradation kinetics
Swovick et al., Mol. Cell. Prot., 2021
Notes: So in this case, we measured that in human primary fibroblasts, HSP90B1 is degraded at 0.21 per day.

## Slide 111: Example Application of LC-MS/MS: Measuring differences in protein expression levels
Swovick et al., Mol. Cell. Prot., 2021
Notes: Additionally, we can measure differences in protein expression levels between various biological conditions.
Here, I measured the protein levels in moue and naked mole rat primary fibroblasts in triplicate and measured the log2 ratio in intensity and the p value of the log2 difference.
And then you can plot those in what’s called a volcano plot. In this particular case, we found that protein involved in ATP production were systematically expressed at higher levels in mouse cells. And we could then verify that difference by performing western blots on a few key proteins.

## Slide 112: Example Application of LC-MS/MS: Reversible post-translational modification (PTM) of amino acids
Ub
Ubiquitination
6-N-Acetyllysine
Phosphoserine
Phosphothreonine
Phosphotyrosine
Notes: Yet another incredibly useful tool for MS/MS is the identification and quantification of post translational modifications. Which, as you can see here, there are many

## Slide 113: Example Application of LC-MS/MS: Reversible post-translational modification (PTM) of amino acids
PTM
1
2
m/z
Δm
1
2
Notes: Each PTM adds a new mass onto the peptide, which results in a mass shift. So through mass spec, we can calculate this difference in shift and determine which modification was located on which part of a protein

## Slide 114: Example Application of LC-MS/MS: Reversible post-translational modification (PTM) of amino acids
Liu et al., PNAS, 2021
Notes: So we can see which phosphorylated proteins are expressed at higher levels between two conditions through a volcano plot again.

## Slide 115: Example Application of LC-MS/MS: Tandem Mass Tags (TMT)
Notes: Finally, the last thing I’ll talk about is something we’ve been doing a lot of over the last few years and that’s tandem-mass tags, or TMT. So this is a method we can use to combine up to 18 different samples together and run them at the same time.
Briefly, the way this works is we take up to 18 of your samples (these can be replicates, some kind of drug treatment, etc) and digest them. Then we add a label to the digests and combine them ad run them o the mass spec.

## Slide 116: Example Application of LC-MS/MS: Tandem Mass Tags (TMT)
Notes: And then for a given peptide, all of the labels have the same m/z on the MS1 spectra. But when we fragment the samples, the labels also fragment at different locations and have different m/z at this second fragmentation. So with this now, we can use the intensity of each label as a proxy for the intensity of a given peptide in each of the 18 samples, thus giving us relative levels of protein abundance between these samples.

## Slide 117: Example Application of LC-MS/MS: Tandem Mass Tags (TMT)
Notes: And as an example of a TMT experiment we did for a customer, they gave us 16 samples; 4 different treatments each with 4 replicates (so we could get some really powerful statistics). And then we can compare the difference intensity between each different condition along with the p value of that differences. And this was all done in one single experiment.
So TMT labeling is a great technology that has allowed for great strides in proteomics over the last 5 years.

## Slide 118: Questions?